# Santa Barbara Pulmonary Associates and Santa Barbara Cottage Hospital Cystic Fibrosis, Bronchiectasis and NTM Program

# Nebulizer Cleaning and Disinfecting Protocol:

# Aerobika OPEP Device & AeroEclipse XL Nebulizer

## Scope

This protocol applies to the cleaning and disinfection of:

- Aerobika® Oscillating Positive Expiratory Pressure (OPEP) Device (Trudell Medical International)

- AeroEclipse® XL Breath Actuated Nebulizer (BAN) (Monaghan Medical Corporation)

Both devices are classified as semi-critical medical devices and require cleaning and disinfection after each use to prevent microbial contamination and reduce the risk of respiratory infection.

---

## General Principles

1. Frequency: Clean both  after each use and disinfect AeroEclipse XL at end of every day or other day. Aerobika may be disinfected weekly.

2. Hand hygiene: Perform hand hygiene before and after handling the device.

3. Water quality: If using a cold (chemical) disinfection method, the final rinse must be performed with sterile water or 0.2 µm filtered water — never tap water. Sterile water rinse is not required after heat disinfection methods.

4. Drying: All components must be completely air-dried before reassembly and storage. Gram-negative bacteria survive better in moist environments.

5. Storage: Store all dried components in a clean, dry, closed container designated solely for that purpose.

6. Replacement: Follow manufacturer-recommended replacement schedules for all components.

---

## SECTION A: Aerobika® OPEP Device

### Step 1 — Disassembly

- Disassemble the Aerobika into its individual components:

- Mouthpiece/cover

- Cartridge housing

- One-way valve

- Pressure indicator (if applicable)

- Remove any attached nebulizer components if used in combination with a nebulizer.

### Step 2 — Cleaning

1. Wash all disassembled parts in warm water with mild liquid dish soap.

2. Gently agitate or use a soft brush to remove any visible residue or secretions.

3. Rinse all parts thoroughly under running water to remove soap residue.

### Step 3 — Disinfection

Select one of the following methods compatible with the manufacturer's instructions:

Heat Methods (preferred — no sterile water rinse required afterward):

- Boiling: Place all parts in continuously boiling water for 5-10 minutes.

- Electric steam sterilizer (e.g., baby bottle sterilizer): Follow sterilizer instructions.

- Dishwasher (top rack): Run a heated dry cycle at ≥70°C (158°F) for at least 30 minutes.

Cold (Chemical) Methods:

- 70%–90% isopropyl or ethyl alcohol: Soak all parts for 5 minutes, then rinse with sterile or 0.2 µm filtered water.

- 3% hydrogen peroxide: Soak all parts for 30 minutes, then rinse with sterile or 0.2 µm filtered water.

Note: The Aerobika device is also compatible with microwave disinfection — place parts in a microwave-safe bowl of water and microwave for 5 minutes. Confirm microwave compatibility per manufacturer IFU.

### Step 4 — Drying and Storage

1. Shake off excess water.

2. Place all parts on a clean, lint-free towel or drying rack.

3. Allow to completely air-dry — do not use a cloth to dry interior surfaces.

4. Once fully dry, reassemble and store in a clean, closed container.

## SECTION B: AeroEclipse® XL Breath Actuated Nebulizer

### Step 1 — Disassembly

- Disconnect the nebulizer from the compressor tubing.

- Disassemble into individual components:

- Nebulizer cup/reservoir

- Top cap/baffle assembly

- Mouthpiece or mask

- T-piece (if applicable)

- Actuator button/mechanism

- Do not disassemble the compressor tubing; replace tubing per manufacturer schedule.

### Step 2 — Cleaning

1. Immediately after each treatment, disassemble the nebulizer.

2. Wash all disassembled parts (except tubing) in warm water with mild liquid dish soap.

3. Use a soft brush to clean the nebulizer cup interior and baffle.

4. Rinse all parts thoroughly under running water.

### Step 3 — Disinfection

Select one of the following methods compatible with the manufacturer's instructions:

Heat Methods (preferred — no sterile water rinse required afterward):

- Boiling: Place all parts in continuously boiling water for 5 minutes.

- Electric steam sterilizer (e.g., baby bottle sterilizer): Follow sterilizer instructions. Studies confirm thermal disinfection with a baby-bottle sterilizer does not alter nebulizer performance over 180 cycles.

- Dishwasher (top rack): Run a heated dry cycle at ≥70°C (158°F) for at least 30 minutes.

Cold (Chemical) Methods:

- 70%–90% isopropyl or ethyl alcohol: Soak all parts for 5 minutes, then rinse with sterile or 0.2 µm filtered water.

- 3% hydrogen peroxide: Soak all parts for 30 minutes, then rinse with sterile or 0.2 µm filtered water.

Note: Microwave disinfection may be used if the device is microwave-safe per manufacturer IFU — place parts in a microwave-safe bowl of water and microwave for 5 minutes.

### Step 4 — Drying and Storage

1. Shake off excess water.

2. Place all parts on a clean, lint-free towel or drying rack.

3. Allow to completely air-dry — do not towel-dry interior surfaces.

4. Once fully dry, reassemble and store in a clean, closed container.

5. Run the compressor briefly (without the nebulizer cup) to dry the interior of the tubing if needed.

## Troubleshooting & Reminders

| Issue | Action |
| --- | --- |
| Visible residue after cleaning | Re-clean with soap and water; use a soft brush |
| Discoloration or warping of parts | Replace the affected component |
| Persistent odor after disinfection | Repeat full cleaning and disinfection cycle; replace if persistent |
| Device not functioning properly | Inspect all parts for damage; replace per manufacturer schedule |
| Tap water used for final rinse after cold disinfection | Repeat disinfection and rinse with sterile/filtered water |

---

## References for Protocol Development

This protocol is based on the Cystic Fibrosis Foundation Infection Prevention and Control (IPC) Guidelines (2014), endorsed by SHEA and APIC, and general infection control principles for semi-critical aerosol delivery devices. Always verify compatibility with the manufacturer's current Instructions for Use (IFU) for each specific device.

1.Device Cleaning and Infection Control in Aerosol Therapy.

Respiratory Care. 2015. O'Malley CA.Review

2.The effects of nebulizer hygiene training on the practices of cystic fibrosis patients and caregivers.

Pediatric Pulmonology. 2021. Yilmaz Yegit C, Ergenekon AP, Mursaloglu HH, et al.

3.A Survey of Cystic Fibrosis Center Recommendations for Nebulizer Cleaning and Disinfection.

Respiratory Care. 2024. Kenkare ZJ, Rearick J, Schramm CM, Murray TS, Collins MS.

4. Nebulizer Cleaning and Disinfection Practices in Families With Cystic Fibrosis: The Relationship Between Attitudes, Practice and Microbe Colonization.

Journal of Cystic Fibrosis : Official Journal of the European Cystic Fibrosis Society. 2019. Murray TS, O'Rourke TK, Feinn R, Drapeau G, Collins MS.

5.Effect of Disinfection Method and Testing Methodology on the Performance of a Breath-Enhanced Jet Nebulizer.

Pharmaceutics. 2023. Agoramurthi K, Berlinski A.

6.Ozone disinfection of home nebulizers effectively kills common cystic fibrosis bacterial pathogens.

Pediatric Pulmonology. 2018. Towle D, Baker V, Schramm C, et al.

Top of Form

Bottom of Form